	       Myung Technology Company

경기도 용인시 처인구 모현면 갈담로 7
Tel# 010-7320-9368  031-726-9368 



          (Chamber Description and specification)
공급하게될 모델은 TPC411입니다만 동일 모델 CATALOG를 구할수 없어서 동일한 사양인 model TPC-412M 에 대한 사양을 참조해 주십시요.
[image: image1.png]Mode T

TPC-212(4) TPC-222(M) TPC-412(M) TPC-42204)

Control sethod |

Method

Teaperature and humidity setting
Control action

Continuous control of sel temperature and hualdity. Program
operatlon. Temperature heat-up or pull-down rate - Hunldity
control (M type)

Direct setting of teaperature and relative hunldlly

PID control. SSR drive system

Power supply 200V AC 41 50/60 Kz 20VAC 1 50/60 He
Total loaded current | 200V AC | 15A 204 | 154 204
2w | Laa 18. SA[  14A 18. 54

Perfornance

Unsaturated control
Teaperature control range
Relative huaidity control range
Pressure contral range
Teaperature and humidity cons-
tancy
Temperature unifornity
Heat up and pressurizing tlze

(al RT+23°C)

+105.0 ~+142.8 °C +105.0 ~+162.2 °C
75~100%RH
0.2 ~ 2.0ke/dG 0.2 ~ 4.0kg/cdG
£0.5 C/* 3%RH

+0.5 °C (at 100%RH) £0.7 °C (at 75%RH)

0 — 2.0k /cdG 0 — 4.0ke/cdG
About | About About About
30 alnutes | 60 minutes 45 oinutes 90 minutes

Saturated control
Temperature control range
Pressure control range
Temperature constancy
Teaperature uniforaity
Heat up and pressurizing tine

(at RT+23C)

+105.0 ~+132.9 °C +105.0 ~+15L.1 C

0.2 ~ 2.0ke/adG 0.2 ~ 4.0e/alG
+05°C
+0.5°C
0 — 2 0ke/adG 0 — 4.0kg/cdG
About About About About
30 minutes | 60 minutes 45 alnutes | 90 mlnutes

Dry and wel-bulb teaperature control

Temperature rise
Temperature control range
Relative hunidlty control range
Heat up and pressurizing time

(a1 RT +23°C)

+105.6 ~+142.0 C | +105.6 ~+162.2 °C

75~95%RH
0.2 — 2.0lg/cdG 0.2 — 4.0kg/cdG
About About Avout About
60 minutes 90 ninutes 75 minutes | 120 minutes

Test process
Temperature control range
Relative humidlty control range
Pressure control range
Tesperature and hualdity cons-
tancy
Teaperature uniforaity

+105.6 ~+1429C | +105.6 ~+162.2 °C
75~9B%RH
0.2 ~ 2.0kg/cdG 0.2 ~ 4.0kg/cdG
0.5 C/*3I%RH

£0.5°C (al 98%RH) -+ 20.7 °C (al T54RH)





[image: image2.png]Mode ! re-mzon | teeemomy | ez | mc

Performance

Dry and wet bulb temperaturs contro

Tesperatre pul|-doun
Tenperature control range +50. 0~+85. 0°C
Temperature anli-down rate from +142.8 C/75%RH | Fron +182.2 C/T5%RH
{(at RT +23°C no-specimen) To +85.0C/85%RH To +85.0°C/85XR 1T
Vin. about 120 ainutes

Hiid pracess
Temperature control range
Relative humidity conlrol range

J Cloth wick 200 hgurs continuows uvse (wilh no specimen)
Construction
Daor Left hinge. Loeking handle
tdoor hook lever only for TPC-212 (M)
Innzr cylinder Specimen shelves and 2-supporls
Control console Teaprature and huaidily indicator, Tine indicator, Key swilch,
Setting buttons, Process indicalor lamps. Trouble alarm lamps
Uthers Exhaust valve, Air inlet valve, Drain filter, Drain valve and
A teat-in puap
[ |
Specimen signal lerminals ICﬂxmlclnr of 12 pins (125V AC/OC 1A)
Pressure gauge Bourdon type

f
Seale -Theakg ~ dg./cxil Seale -ToenHg ~10kz./ ot

r

Pressure vessa! overheal pratector, Boii-dry protectar, fver-
pressure protector. Power [ailure safely mechanisa Leakage
breaker for power supply. Safety valve. Temp. sensar discon-
nection safety mechanisn. Agilator fan slop eechaaisa. Wel-
oulh wick drying protector mechanisn
Control connector for external For external alarn signal
alam and spacizen power supply Function: Contact outpul ON when in trovble status
For specimen power supply signal

Funclion: Contact output DN whea in test stalus

Safely mechanisn

Auto star: timer Digital setting, indication
T
Dimensions |
Tesl area volune (2) 2 55 2% 55
Inside diameter of test ares 20.5x3 38.5x45 20.5%30 38.5%45
(@ xL cn)
Dutside dianeter  (WXIIXD crj 51x 13580 | 6lx M7x95| 5lx 136xB0| Blx 14795
Weight (@) ‘ About 117 | Abeut 180 Aout 127 About 200





[image: image3.png]


 [image: image4.png]



[image: image5.png]


[image: image6.png]ATURE AND HOWHDITY CHAMBER

MO TPC-411
SYSTEM BALANCED PRESS. — TEMP. AND

HUMID. CONTROL
POWER. SOURCE AC V 14 50/60Hz MAX. A
TEMP. RANGE +105~+162.2°C
» HUMID. RANGE 75 ~ 100 2 RH ®
" PRESSURE 4.0 kg/cm2G -
HEATER 0.5 kW HUMID. HEATER  1.5kW
CHAMBER. HEATER 0.062 kW
INSIDE. DIMENSIONS $ 295 XL 30cm
WEIGHT 127 kg
DATE SERIAL NO.

0saka - TABAI ESPEC CORP. AN




 [image: image7.png]» LV —-& 7 - - - a .
HOUR METER




  [image: image8.png]


  
온도, 습도와 압력 기록계(Digital recorder 가 필요할 경우 대체 가능함, 온도 습도 analog output 있음)

Page 4/페이지 1 / 2

